## Slide 1
1
1
1
0
0
0
Pressure sensor
Sample IN
Pump OUT
V1
V1
V2
In-line filter
Pump & V4
1=ON, 0=OFF
States | Description | V1 | V2 | V3 | V4
Gasx trig (LV) | Flush cell for T_flush (goto Evac to P_meas) | 1 | 1 | 0 | 1
Evac to P_meas | Choked evac to P<P_cell_evac to meas or wait Evac_TO (goto Measure P) | 0 | 1 | 1 | 1
Measure P | Cell open to pressure sensor for T_pressure (if “Measure_pressure” goto Idle, else goto Measure) | 0 | 1 | 1 | 0
Measure | Wait for T_meas (if “GasX_on” goto Gasx trig, else goto Idle) | 0 | 0 | 0 | 0
Idle | Wait for GasX trigger (goto Gasx trig) or shutdown or “Measure_pressure” (goto Measure P) | 0 | 0 | 0 | 0
Valve logic:
0: default (down)
1: active (up)
“ “: common
1
0
Adj. choke
Currently used parameters:
T_flush
P_cell_evac to meas
T_pressure
T_meas
Evac_TO
GasX_on
Measure_pressure
V4
V3
1
0